| SP1000UV/SP1000/SP2000UV/SP2000 UV/Visible Spectrophotometer |  |  |  |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| Instruction Manual |  |  |  |  |  |  |  |  |
| PN : 435057 |  |  |  |  |  |  | Version : 1.0.1 |  |

Contents

Safety Information	1

Symbols and Conventions	1

Specifications	2

Overview	2

Description of Buttons and Switches	3

Getting Started	3

General Operating Instructions	4

Calibration and System Settings	7

Performance Verification	11

Measurement	14

Troubleshooting	19

Repair and Maintenance	20

Warranty	23

Compliance with Local Laws	23

## Safety Information

Please read this manual thoroughly before operating the unit.
Follow all safety and operating guidelines to ensure proper use of the instrument.

Do not open the device.

Always disconnect the device from the main power supply before performing any maintenance or replacing fuses.

Danger: High voltage inside the device. Risk of electric shock.

Do not use the device if it is damaged, especially if the main power cable is frayed, cut, or defective.

Repairs must only be performed by authorized service technicians from your local MPD Scientific office or approved service partners.

Ensure the device is connected to a properly grounded power outlet.

Using the equipment in a manner not specified by the manufacturer may impair the protective features of the instrument.

Prevent liquids from entering the device.

Do not operate the device in hazardous or potentially explosive environments.

Package Contents

| Description | Quantity |
| --- | --- |
| Spectrophotometer | 1PC |
| Glass Cuvette | 4PCS |
| Quartz Cuvette | 2PCS |
| Power Cord | 1PC |
| Instruction Manual | 1PC |
| Dust Cover | 1PC |

Unpacking

After opening the package, carefully check the contents against the packing list. If any items are missing or damaged, please contact your local MPD Scientific office or an authorized service partner immediately.

Installation

Placement
Place the instrument carefully on a stable, level surface to ensure proper operation and to prevent vibration or accidental movement.

Connecting the Power Cord
Before connecting, make sure the instrument’s power switch is turned off.
Insert the female end of the power cord into the instrument’s power inlet, then plug the male end into a properly grounded power outlet.

## Symbols and Conventions

The following chart is an illustrated glossary of the symbols that are used in this manual.

|  | CAUTION Indicates a potential risk and alerts the user to proceed with caution. |
| --- | --- |
|  | HIGH VOLTAGE Warns of the presence of high voltage. Follow all safety procedures to avoid electric shock. |
|  | HOT SURFACE Indicates a risk of burns due to hot surfaces. Avoid contact during or after operation. |

## Specifications

| Model | SP1000 | SP2000 | SP1000UV | SP2000UV |
| --- | --- | --- | --- | --- |
| Optical system | Single beam，1200 l/mm grating |  |  |  |
| Light source | Tungsten lamp |  | Tungsten lamp，Deuterium lamp |  |
| Detector | Silicon photodiode |  |  |  |
| Spectral bandwidth | 4 nm | 2 nm | 4 nm | 2 nm |
| Wavelength range | 320~1050 nm |  | 195~1050 nm |  |
| Wavelength accuracy | ±0.8 nm |  |  |  |
| Wavelength repeatability | ≤0.5 nm |  |  |  |
| Wavelength resolution | 0.1 nm |  |  |  |
| Wavelength swing speed | 6000 nm/s |  |  |  |
| Photometric range | -0.3~3 A，0~200 %T，0~9999.9 C |  |  |  |
| Photometric accuracy | ±0.002 A（0~0.5 A），±0.004 A（0.5~1 A）, ±0.5 %T（0~100 %T） |  |  |  |
| Photometric repeatability | ±0.001 A（0~0.5 A），±0.002 A（0.5~1 A）, ±0.2 %T（0~100 %T） |  |  |  |
| Stray light | ≤0.2 %T（360 nm） |  | ≤0.2 %T（220 nm， 360 nm） |  |
| Noise | ≤0.001 A @ 0 A，500 nm，≤0.002 A @ 1 A，500 nm，≤0.004 A @ 2 A，500 nm |  |  |  |
| Sample charger | 10 mm 4-cell manual charger |  |  |  |
| Display | 5 inch TFT color touch screen |  |  |  |
| Storage | 236 KB (built-in), supports unlimited external expansion (USB memory) |  |  |  |
| Interface | RS232 serial port × 1 (printer), USB-A × 1 (USB memory/USB printer), USB-B × 1 (computer) |  |  |  |
| Power supply | 100~240 V AC，50~60 Hz，75 W |  | 100~240 V AC，50~60 Hz，120 W |  |
| Size | 456（W）×360（D）×185（H） mm |  |  |  |
| Weight | 10.5 kg |  | 10.7 kg |  |

## Overview

The SP1000 / SP2000 / SP1000UV / SP2000UV Spectrophotometers are versatile analytical instruments designed for precise photometric and wavelength measurements across a wide range of laboratory and industrial applications. These instruments are commonly used in chemistry, pharmaceuticals, biochemistry, metallurgy, light industry, textiles, materials science, environmental testing, medical research, and education, particularly in quality control and research laboratories.

## Description of Buttons and Switches

Front View

Right View

Rear View

## Getting Started
The following steps outline the basic operation of the instrument.

Power On and Self-Check

- Switch on the power.
After the instrument is powered on, it will automatically perform a self-check sequence to ensure proper operation.

- Self-check procedures include:

  - Lamp ignition

  - Filter disc positioning

  - Automatic sample holder positioning (if installed)

  - Dark current acquisition

  - Wavelength calibration

  - Energy verification

  - System baseline check

Once the self-check process is complete, the instrument is ready for measurement.

## General Operating Instructions

Touchscreen Operation

The instrument is equipped with a full-color touchscreen that allows easy operation through direct touch.
To select an option, gently press the desired area of the screen using your fingertip, fingernail, or a stylus.
Avoid using sharp objects such as ballpoint pens or metal tools, as these may damage the screen surface.

Selecting an Application

From the Main Interface, touch the corresponding icon to open and begin the desired application.

|  | Photometry Measure the absorbance or transmittance of the sample. |
| --- | --- |
|  | Quantitation Establish the standard curve and measure the concentration of the sample. |
|  | Spectrum Scan the sample in a wavelength range. |
|  | File Manage files stored in the instrument or USB disk. |
|  | System calibration and setup. |
|  | Performance verification Verify the performance of the instrument. |

Basic Operation

|  | Home Back to main interface. |
| --- | --- |
|  | Return to the previous interface. |
| / | Page Up/Down Go to previous/next page. |

Measurement Results Operation

|  | Open result(s) from internal/USB memory. |
| --- | --- |
|  | Save result(s) to internal/USB memory. |
|  | Print result(s). |
|  | Delete selected result(s). |

Rename, Print and Delete Results

| Rename a Sample: | In the list interface, tap the Name field of the desired sample. Enter a new sample name (up to 8 characters) using the on-screen keyboard. |
| --- | --- |
| Print the Measurement Report: | In the list interface, tap the Print icon to generate and print the measurement report. |
| Delete sample(s): | Select the desired sample(s) by checking the corresponding box. Tap the Delete icon to remove the selected record(s). |

Open Results

| Open: | In the list interface, press the icon . Select the storage location by tapping Internal Memory or USB Memory, depending on where the file is saved. From the file list, choose the desired file, then tap Open to load it. |
| --- | --- |

Save Results

| Save: | In the list interface, press the icon Save. Press the icon / to select the Internal/USB where the file is saved. Type in the file name, press the button Save. |
| --- | --- |

Files Operation

|  | Internal Memory Internal memory of the spectrophotometer. |
| --- | --- |
|  | USB Memory USB extended mass memory. |
|  | Copy the selected file(s) from internal /USB memory to USB/internal memory. |
|  | Export csv Export file(s) to *.csv format |
|  | Export txt Export file(s) to *.txt format |
|  | Delete the selected file(s). |

Rename, Import, Export and Delete Files

| Rename a File: | In the file management interface, tap the Name field of the desired file. Enter a new file name (up to 8 characters) using the on-screen keyboard. |
| --- | --- |
| Copy File(s) From/To Internal Memory/USB Memory: | In the File management interface, press the Check Box, press the button （USB Disk required）. |
| Export File(s) To *.csv Format | In the File management interface, press the Check Box, press the button （USB Disk required）. |
| Export File(s) To *.txt Format | In the File management interface, press the Check Box, press the button （USB Disk required）. |
| Delete File(s): | In the File management interface, press the Check Box, and press the icon . |

## Calibration and System Settings

Select the icon  in the main interface. It will display options to calibrate the system and configure basic instrument settings.

Calibration

|  | Calibrate: Begin calibration sequence. |
| --- | --- |

Select Tab Calibration in the System interface. Ensure that all samples are removed from the measurement channel. Close the sample chamber cover securely. Select one of the following calibration options to perform the desired calibration procedure: Wavelength Calibration, System Baseline Calibration, Dark Current.

Confirm your selection by pressing the  icon.

Settings of Light Source

|  | Tungsten Lamp Reset |
| --- | --- |
|  | Deuterium Lamp Reset |

Select Tab Light source in the System interface to display light source information.

Visible models

UV/VISIBLE models

| On/Off lamp: | Press the icon to turn on/off the Tungsten lamp/Deuterium lamp. |
| --- | --- |
| Change the lamp switching point: | Press the value of lamp switching point. Input the new value. |
| Reset the lamp usage: | Press the icon / to reset the Tungsten lamp/Deuterium lamp usage time. |

Important information If only one of the light sources is used for a long period of time, please turn off another light source to save energy. If the lamp switching point is changed, the system baseline must be recalibrated.

Edit Clock

|  | Accept the new value. |
| --- | --- |

Select Tab Clock in the System interface. Press the value of year, month, date, hour, minute or second to change. Press the icon  to accept new value.

Memory Management

|  | Format Internal Memory. |
| --- | --- |
|  | Format USB Memory storage. |

Select Tab Memory in the System interface. The use of the internal and USB memory (If inserted) are shown. Press the icon  /  to format internal memory/USB memory.

Language Selection

|  | Accept the new language. |
| --- | --- |

Select Tab Language in the System interface. Select a language, press the icon  to change.

General Options

Select the Tab General Options in the System interface.

| Beep: | Press the icon to turn on/off the beep. |
| --- | --- |
| Brightness: | Press the icon to decrease/increase the brightness of the LCD display. |
| Close display: | Press the icon to turn the display on/off. When turned on, the display will close automatically if there is no operation for 30 minutes. |
| Select sample holder | If your instrument is equipped with an automatic sample cell holder, you must configure the holder type before first use. Press to select the installed holder model: Auto Five-Cell Holder (AC-5) or Auto Eight-Cell Holder (AC-8). |

Restore Defaults

|  | Restore the parameters to factory settings. |
| --- | --- |

Select the Tab Restore defaults in the System interface. Select an item, press the icon  to restore.

## Performance Verification

Select the icon  in the main interface. This will display options to verify the performance of the instrument.

Important information Before verifying instrument performance, allow the instrument to preheat for 30 minutes. After preheating, re-measure the dark current to ensure accurate performance verification.

Verifying Wavelength Accuracy and Wavelength Repeatability

Select the Tab Wavelength Accuracy in the Performance verification interface.

| Standard Sample: | Holmium oxide solution or equivalent filter |
| --- | --- |
| Measurement: | Prepare for Measurement Remove any sample from the measurement channel. Close the sample chamber cover. Press the wavelength value field, enter the desired measurement wavelength, and press the Zero button. Measure the Standard Sample Place the standard sample in the measurement channel and press the Measure button. Repeat Measurements Repeat Step 2 three times. - The difference between the average of the three readings and the standard value represents the single-point wavelength tolerance. - The difference between the maximum and minimum values among the three readings represents the single-point wavelength repeatability. Verify Across Wavelengths Repeat Steps 1–3 for each selected wavelength point. - The maximum value among all single-point wavelength tolerances represents the instrument’s wavelength accuracy. - The maximum value among all single-point wavelength repeatabilities represents the instrument’s wavelength repeatability. |

Verifying Photometric Accuracy and Photometric Repeatability

Select Tab Photometric accuracy in the Performance verification interface.

| Standard Sample: | NIST 930D or equivalent filter |
| --- | --- |
| Measurement: | Prepare for Measurement Remove any sample from the measurement channel and close the sample chamber cover. Press the wavelength value field, enter the desired measurement wavelength, and press the Zero button. Measure the Standard Sample Place the standard sample in the measurement channel and press the Measure button. Repeat Measurements Repeat Step 2 three times. - The difference between the average of the three measurements and the standard value represents the single-point photometric tolerance. - The difference between the maximum and minimum values among the three measurements represents the single-point photometric repeatability. Verify Across Wavelengths Repeat Steps 1–3 for each selected wavelength point. - The maximum value among all single-point photometric tolerances represents the instrument’s photometric accuracy. - The maximum value among all single-point photometric repeatabilities represents the instrument’s photometric repeatability. |

Verifying Stray Light

Select Tab Stray light in the Performance verification interface.

| Standard Sample: | 10g/L NaI solution or equivalent filter (220nm, UV/VISIBLE models only), 50g/L NaNo2 solution or equivalent filter (340 or 360nm) |
| --- | --- |
| Measurement: | Remove any sample from the measurement channel and close the sample chamber cover. Press the wavelength value field, enter the desired measurement wavelength, and press the Zero button. Measure the Standard Sample Place the standard sample in the measurement channel and press the Measure button. |

Verifying Noise

Select Tab Noise(0A) in the Performance verification interface.

| Standard Sample: | None |
| --- | --- |
| Measurement: | Remove any sample from the measurement channel and close the sample chamber cover. Press the wavelength value field, enter the desired measurement wavelength, and press the Zero button. Measure the Standard Sample Place the standard sample in the measurement channel and press the Measure button. |

Verifying Dark Noise

Select Tab Noise(0%T) in the Performance verification interface.

| Standard Sample: | Block |
| --- | --- |
| Measurement: | Remove any sample from the measurement channel and close the sample chamber cover. Press the wavelength value field, enter the desired measurement wavelength, and press the Zero button. Measure Dark Noise Place the light-blocking device (Block) in the measurement channel and press the Measure button. The displayed result represents the dark noise at the selected wavelength. |

Verifying Stability

Select Tab Stability in the Performance verification interface.

| Standard Sample: | None |
| --- | --- |
| Measurement: | Remove any sample from the measurement channel and close the sample chamber cover. Press the wavelength value field, enter the wavelength 500, press the Zero button. Press the button Measure, the result is the noise of 500nm. |

Verifying Bandwidth

Select Tab Bandwidth in the Performance verification interface.

| Standard Sample: | Low pressure quartz mercury lamp |
| --- | --- |
| Measurement: | Open the lamp cover, put the low-pressure quartz mercury lamp into the lamp seat, and turn it on. Remove any sample from the measurement channel, close the sample chamber cover, press the wavelength value, type in the wavelength 546.1; Press the button Measure, the result is the bandwidth. |

## Measurement

Important Guidelines

Reagents and dilution buffers may cause burns or other health hazards. Samples such as nucleic acids, proteins, and bacterial cultures can be infectious and may cause serious damage to health. During sample preparation, measurement, maintenance, and cleaning, always follow local laboratory safety regulations. Wear appropriate personal protective equipment, including lab coats, gloves, and eye protection, and use disinfectants as required. Dispose of measurement solutions, cleaning agents, and disinfectants in accordance with local laboratory regulations.

Cuvette Inspection

Ensure that all cuvettes are clean, transparent, and free of sample residue. Only quartz cuvettes are permitted for measurements in the UV wavelength range.

Photometry

Photometry mode is used to measure the absorbance or transmissivity of the sample.

Within the Main interface, press the icon  to start a Photometry application.

|  | Mode Switch measurement mode to %T, Abs or Energy. |
| --- | --- |
|  | Wavelength Set measurement wavelength. |
|  | Zero Do 0Abs/100%T. |
|  | Read Measure sample and record the result. |
|  | List View the result(s) list. |
|  | Increase/Decrease the gain of signal. Only for Energy mode. |

Press the icon  to switch to the measurement mode.

| Abs | Measure absorbance value of the sample(s). |
| --- | --- |
| %T | Measure transmittance value of the sample(s). |
| E | Measure energy value of the sample(s). |

Press the icon  to set wavelength, key in the measurement wavelength.

Put the reference in the measurement channel, press the icon  to zero.

Put the sample in the measurement channel, press the icon  to measure a sample and record the result.

Press the icon  to browse the result(s).

Quantitation

Quantitation mode is used to measure the concentration of the sample(s).

Within the Main interface, press the icon  to start a Quantitation application.

Establish Method

Within the Quantitation interface, press the button Establish method.

| Measurement | A = A₁ Absorbance (A) is equal to the measured absorbance at wavelength 1 (A₁). A = A₁ − m × A₂ Absorbance (A) is equal to the difference between the absorbance at wavelength 1 (A₁) and the absorbance at wavelength 2 (A₂) multiplied by a coefficient (m). A = A₁ / A₂ Absorbance (A) is equal to the ratio of the absorbance at wavelength 1 (A₁) to the absorbance at wavelength 2 (A₂). |
| --- | --- |
| Wavelength 1 | Measurement wavelength 1 |
| Wavelength 2 | Measurement wavelength 2 |
| Fitting | LIN-0: Linear to zero LIN: Linear. QUA: Quadratic. |
| Unit | - (No Unit), %, ppm, ppb, g/L, mg/L, μg/L, ng/L, g/dL, mg/dL, μg/dL, mg/mL, μg/mL, ng/mL, μg/μL, ng/μL, mol/L, mmol/L, IU, Custom(User input, Up to 8 characters). |
| Calibration | Coe K: Input equation coefficient. Std M: Measure standard sample(s) Std I: Input standard sample(s) |
| Standard quantity | Standard sample number (Up to 10) |

Press the item to configure the measurement parameters.
After all parameters are set, press the Next button to begin establishing the standard curve.

- If the Calibration item is set to Coe K, Std M, or Std I, refer to sections 2.3.1, 2.3.2, or 2.3.3, respectively.

Enter the equation coefficients to establish the standard curve.
Input coefficients K₀ through K₃, then press the Next button to continue.

Measure standard sample to establish standard curve

Put the reference in the measurement channel, press the button Zero to do zero.

Put the 1# standard sample in the measurement channel, press the button Read to measure.

Repeat step 3.3 to measure other of the standard samples.

Press the item to input concentration of standard samples, press the button Next.

Establishing the Standard Curve

Place the standard sample in the measurement channel to establish the standard curve.

Press the Abs and Conc items to enter the absorbance and concentration of each standard sample, then press the Next button.

After completing the standard curve setup, press Save to save the method.
Press Measure to apply the new method and enter the measurement interface.
Press Finish to exit the setup.

Measure Sample

Quantitation interface, press the button Measure sample.

|  | Method Select Measurement Method. |
| --- | --- |
|  | Zero 0Abs/100%T. |
|  | Read Measure the sample and record the result. |
|  | List View the result(s) list. |

Press the icon  to select method.

Press the Open button to load a measurement method stored in the internal memory or USB drive.
Press Measure to apply the selected method and return to the measurement interface. Put the reference in the measurement channel, press the icon  to do zero.

Put the sample in the measurement channel, press the icon  to measure a sample and record the result.

Press the icon  to browse the result(s).

## Troubleshooting

Review the information in the table below to solve operating problems.

| Problem | Cause | Solution |
| --- | --- | --- |
| Power on, no response | Power cord connection is not reliable | Improve connectivity |
|  | Fuse burning | Replace fuse |
| Measurement uncertainty | Sample is not Stable | Improve the sample |
|  | Glass cuvettes used in UV Range | Use quartz cuvettes |
|  | The concentration of sample is too high | Diluted sample |
|  | Power Supply Voltage Low or not Stable | Improve the Power Supply |
|  | Lamp damage or lamp life maturity | Replace lamp |
| Dark Current Error when self-check | The lid of the compartment is open during self-check | Close the lid, restart |
| System Calibrate Failed | Something block the Light path | Remove it, calibrate again |
| Measurements inaccurate | Cuvettes were contaminated | Clean cuvettes |
|  | Samples were contaminated | Improve samples |
|  | Worse matching of the cuvettes | Improve the matching of the cuvettes |
|  | Dark current error | Resample dark current |

## Repair and Maintenance

Check the Sample Compartment
After each measurement, promptly remove the cuvettes containing sample solutions from the compartment. If left inside, evaporation may cause condensation and lead to mold formation on the optical mirrors.
Pay special attention when using corrosive or highly volatile samples. Immediately wipe off any spilled or residual liquid inside the compartment using a soft, lint-free cloth.

Clean the Instrument Surface
The instrument cover is coated with protective paint. Wipe any spills or drips from the surface immediately using a damp cloth.
Do not use organic solvents or abrasive cleaners, as these may damage the finish. Clean any dirt or residue from the cover regularly to maintain the instrument’s appearance and durability.

Clean the Cuvettes
After each measurement or whenever the sample solution is changed, carefully clean the cuvettes. Residual liquid or contamination on the cuvette surfaces can cause measurement errors. Ensure the cuvettes are completely dry before reuse or storage.

Spare Parts Replacement

Replace the fuse

|  | Danger! Be sure to switch off the power and unplug the socket before replacement! |
| --- | --- |

Tools preparation

Prepare a 3×75 Flat Blade screwdriver.

Switch Off the power supply

Switch off the power supply and unplug the socket.

Take out the Fuse Seat

Take out the Fuse Seat

Push the fuse case by using the screwdriver, and turn it counterclockwise, the fuse seat will pop out when released.

Replace a new fuse

Pick out the spare fuse (3.15A/250V) and replace it.

Reset the fuse seat

Replace the fuse seat in the power socket. Push the fuse case by using the screwdriver, and turn it clockwise, the fuse seat will be locked when released.

Switch on the power

Plug the socket and switch on the power.

Replace lamps

|  | Hot! Wait at least 20 minutes after turning off the power before opening the lamp chamber to prevent burns. |
| --- | --- |

Tools preparation

Prepare a 6×150mm Flat Blade screwdriver and a pair of glove.

Power Off

Switch off the power supply and unplug the socket.

Open the cover

Loosen the two indicated screws and remove the lamp cover.

Replace the D2 lamp

If your spectrophotometer is a visible model, please skip to Step 5.

Unplug connector No. 2. Unscrew the two screws on the D₂ lamp flange (No. 1) and remove the D₂ lamp. Put on cotton gloves and install a new lamp. Secure it with the two screws, then reconnect the connector.

Replacing the W Lamp
The tungsten lamp is supplied with a blue-grey silicone coating from the manufacturer. This coating is only for transport protection and will be removed during the first lamp replacement.
Remove the defective tungsten lamp while wearing cotton gloves. Insert the new lamp fully into the lamp seat, ensuring that the filament is aligned in the same direction as the old one.

Adjust the Position of the W Lamp

Turn on the power (the switch mirror should be set to the indicated position). Observe the entrance spot—it should be centered in the entrance hole.
If the light spot deviates to the left or right, loosen the two screws and adjust the lamp seat in the corresponding direction until the spot is centered in the slot. Once properly aligned, tighten the screws to secure the lamp seat.

Finish

Reset the cover of the light chamber and fix the screws. Reset the cover of the lamp room and fix the screws.

## Warranty

We warrant that this product is free from defects in materials and workmanship for a period of one (1) year from the date of delivery. If a defect is found during this period, we will, at our discretion and expense, repair the product, replace it, or refund the original purchase price, provided the product is returned within the warranty period.

This warranty does not apply to products that have been damaged as a result of accident, abuse, misuse, improper application, or normal wear and tear. Failure to perform the required maintenance or inspection services in accordance with the operating manual or applicable local regulations will render this warranty void, except where the defect is proven to be unrelated to such non-performance.

Products returned under warranty must be properly insured by the customer against loss or damage during shipment. This warranty is limited to the remedies specified above. IT IS EXPRESSLY AGREED THAT THIS WARRANTY IS IN LIEU OF ALL OTHER WARRANTIES, WHETHER EXPRESS OR IMPLIED, INCLUDING ANY WARRANTY OF MERCHANTABILITY OR FITNESS FOR A PARTICULAR PURPOSE.

# Compliance with Local Laws

The customer is responsible for applying for and obtaining the necessary regulatory approvals or other authorizations necessary to run or use the Product in its local environment. WE will not be held liable for any related omission or for not obtaining the required approval or authorization, unless any refusal is due to a defect of the product.

Equipment Disposal

This equipment is marked with the crossed-out wheeled bin symbol, indicating that it must not be disposed of with unsorted municipal waste.

At the end of its service life, it is your responsibility to ensure proper disposal by delivering the equipment to an authorized collection or recycling facility.

If the equipment has been exposed to biological, chemical, or radiological materials, it must be decontaminated prior to disposal to protect the health and safety of personnel handling recycling and waste processing.

For information on authorized collection points or recycling options, please contact your local distributor or dealer from whom the equipment was purchased.

By disposing of this equipment responsibly, you are helping to conserve natural resources and protect human health and the environment.

Thank you for your cooperation.

609.642.6840 | info@mpdscientific.com

www.mpdscientific.com

Copyright ©2025 MPD Scientific Inc.